## Holding Registers
Spectro::lyzer Holding Registers
Read Command = 0x03
Write Command = 0x06 for 1 register, 0x10 for multiple registers
Group | Tag Name | Register | Type | Length | R/W | Description | Address | Notes
Public Device Configuration | uiAddress | 0x000 | unit16 | 1 | RW | Serial mode: Modbus address of device TCP mode: Portnumber of device | 0
| eCommMode | 0x0001 | enum | 1 | RW | Shows the mode of the device 0 ... Modbus RTU 1 ... Modbus serial ASCII 2 ... Modbus TCP | 1
| eBaudrate | 0x0002 | enum | 1 | RW | Serial mode: baudrate 0 ... 9600 baud 1 ... 19200 baud (default) 2 ... 38400 baud TCP mode: 0 (readonly) | 2
| eParity | 0x0003 | enum | 1 | RW | Serial mode: parity 0 ... no parity 1 ... even parity 2 ... odd parity (default) TCP mode: 0 (readonly) | 3
| eChangeSettings | 0x0004 | enum | 1 | RW | General changes on all settings (e.g. reset all settings to default) | 4
| pDeviceConfigPrivate | 0x0005 | pointer | 1 | R | Pointer to startaddress of device configuration private | 5 | Value incorrect for ana::gate
| abDeviceLocation | 0x0006 | char[12] | 6 | RW | Installation location (s::canpoint) of the device, filled with spaces | 6
| eCleanMode | 0x000C | enum | 1 | RW | Cleaning mode configuration: 0 ... no cleaning supported, manual OFF 1 ... manual ON 2 ... automatic cleaning | 12 | Spec does not respond - does not give correct value or write the correct value. No data in ana::gate
| uiCleanInterval | 0x000D | unit16 | 1 | RW | Cleaning interval 0 ... automatic disabled | 13 | Spec does not respond - does not give correct value or write the correct value. No data in ana::gate
| uiCleanDuration | 0x000E | unit16 | 1 | RW | Cleaning duration in seconds | 14 | No data in ana::gate
| uiCleanWait | 0x000F | unit16 | 1 | RW | Waiting time between end of cleaning and start of measurement | 15 | No data in ana::gate
| tSystemTime | 0x0010 | timestamp (TAI64N) | 6 | RW | Current system time | 16 | No data in ana::gate
| uiMeasInterval | 0x0016 | unit16 | 1 | RW | Measurement interval in sec. 0 ... as fast as possible | 22 | No data in ana::gate
| eLoggingMode | 0x0017 | - | 1 | RW | Spec logging mode 0 … logging mode 1 … online mode | 23 | No data in ana::gate
| uiLogInterval | 0x0018 | unit16 | 1 | RW | Logging interval for data logger in minutes 0 ... no datalogger active | 24 | No data in ana::gate
| uiNLogResultsTotal | 0x0019 | unit16 | 1 | RW | Available number of logged results in datalogger (since last clearing) | 25 | Spec does not respond - does not give correct value or write the correct value
| uiIndexLogResult | 0x001A | unit16 | 1 | RW | Index device status public+ private & parameter results from logger storage to Modbus registers. If no stored results are available results are NaN, Device status bit3 is set. | 26 | Spec does not respond - does not give correct value or write the correct value
| Reserved | 0x001B | - | 93 | R |  | 27 | Seems to be entirely "0"s, with the exception of values in #31, #33 and #35 (from spectro::lyzer only)
Parameter 1 Description | abP1Name | 0x0078 | char[8] | 4 | RW | Name of parameter 1 (filled with spaces) | 120
| abP1Unit | 0x007C | char[8] | 4 | RW | Unit of parameter 1 (filled with spaces) | 124
| xP1UpperLimit | 0x0080 | float32 | 2 | RW | Upper measuring range of parameter 1 | 128
| xP1LowerLimit | 0x0082 | float32 | 2 | RW | Lower measuring range of parameter 1 | 130
| ??? | 0x0084 | ? | 2 |  |  | 132 | 0x8000 0x0000
| fCalibrCoeff | 0x0086 | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 134
| ??? | 0x008E | float32 | 4 |  |  | 142 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0092 | int16 | 1 |  |  | 146 | 0x0003 or 0x0000
| ??? | 0x0093 | int16 | 1 |  |  | 147 | 0x0002
| ??? | 0x0094 | float32 | 8 |  |  | 148 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x009C |  | 84 |  |  | 156 | All zeros
Parameter 2 Description | abP2Name | 0x00F0 | char[8] | 4 | RW | Name of parameter 2 (filled with spaces) | 240
| abP2Unit | 0x00F4 | char[8] | 4 | RW | Unit of parameter 2 (filled with spaces) | 244
| xP2UpperLimit | 0x00F8 | float32 | 2 | RW | Upper measuring range of Parameter 2 | 248
| xP2LowerLimit | 0x00FA | float32 | 2 | RW | Lower measuring range of Parameter 2 | 250
| ??? | 0x00FC | ? | 2 |  |  | 252 | 0x8400 0x0000
| ??? | 0x00FE | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 254
| ??? | 0x0106 | float32 | 4 |  |  | 262 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x010A | int16 | 1 |  |  | 266 | 0x0003 or 0x0000
| ??? | 0x010B | int16 | 1 |  |  |  | 0x0002
| ??? | 0x010C | float32 | 8 |  |  | 268 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0114 |  | 84 |  |  | 276 | All zeros
Parameter 3 Description | abP3Name | 0x0168 | char[8] | 4 | RW | Name of Parameter 3 (filled with spaces) | 360
| abP3Unit | 0x016C | char[8] | 4 | RW | Unit of Parameter 3 (filled with spaces) | 364
| xP3UpperLimit | 0x0170 | float32 | 2 | RW | Upper measuring range of Parameter 3 | 368
| xP3LowerLimit | 0x0172 | float32 | 2 | RW | Lower measuring range of Parameter 3 | 370
| ??? | 0x0174 | ? | 2 |  |  | 372 | 0x8800 0x0000
| ??? | 0x0176 | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 374
| ??? | 0x017E | float32 | 4 |  |  | 382 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0182 | int16 | 1 |  |  | 386 | 0x0003 or 0x0000
| ??? | 0x0183 | int16 | 1 |  |  | 386 | 0x0002
| ??? | 0x0184 | float32 | 8 |  |  | 388 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x018C |  | 84 |  |  | 396 | All zeros
Parameter 4 Description | abP4Name | 0x01E0 | char[8] | 4 | RW | Name of Parameter 4 (filled with spaces) | 480
| abP4Unit | 0x01E4 | char[8] | 4 | RW | Unit of Parameter 4 (filled with spaces) | 484
| xP4UpperLimit | 0x01E8 | float32 | 2 | RW | Upper measuring range of Parameter 4 | 488
| xP4LowerLimit | 0x01EA | float32 | 2 | RW | Lower measuring range of Parameter 4 | 490
| ??? | 0x01EC | ? | 2 |  |  | 492 | 0x8C00 0x0000
| ??? | 0x01EE | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 494
| ??? | 0x01F6 | float32 | 4 |  |  | 502 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x01FA | int16 | 1 |  |  | 506 | 0x0003 or 0x0000
| ??? | 0x01FB | int16 | 1 |  |  | 507 | 0x0002
| ??? | 0x01FC | float32 | 8 |  |  | 508 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0204 |  | 84 |  |  | 516 | All zeros
Parameter 5 Description | abP5Name | 0x0258 | char[8] | 4 | RW | Name of Parameter 5 (filled with spaces) | 600
| abP5Unit | 0x025C | char[8] | 4 | RW | Unit of Parameter 5 (filled with spaces) | 604
| xP5UpperLimit | 0x0260 | float32 | 2 | RW | Upper measuring range of Parameter 5 | 608
| xP5LowerLimit | 0x0262 | float32 | 2 | RW | Lower measuring range of Parameter 5 | 610
| ??? | 0x0264 | ? | 2 |  |  | 612 | 0x9000 0x0000
| ??? | 0x0266 | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 614
| ??? | 0x026E | float32 | 4 |  |  | 622 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0272 | int16 | 1 |  |  | 626 | 0x0003 or 0x0000
| ??? | 0x0273 | int16 | 1 |  |  | 627 | 0x0002
| ??? | 0x0274 | float32 | 8 |  |  | 628 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x027C |  | 84 |  |  | 636 | All zeros
Parameter 6 Description | abP6Name | 0x02D0 | char[8] | 4 | RW | Name of Parameter 6 (filled with spaces) | 720
| abP6Unit | 0x02D4 | char[8] | 4 | RW | Unit of Parameter 6 (filled with spaces) | 724
| xP6UpperLimit | 0x02D8 | float32 | 2 | RW | Upper measuring range of Parameter 6 | 728
| xP6LowerLimit | 0x02DA | float32 | 2 | RW | Lower measuring range of Parameter 6 | 730
| ??? | 0x02DC | ? | 2 |  |  | 732 | 0x9400 0x0000
| ??? | 0x02DE | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 734
| ??? | 0x02E6 | float32 | 4 |  |  | 742 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x02EA | int16 | 1 |  |  | 746 | 0x0003
| ??? | 0x02EB | int16 | 1 |  |  | 746 | 0x0002
| ??? | 0x02EC | float32 | 8 |  |  | 748 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x02F4 |  | 84 |  |  | 756 | All zeros
Parameter 7 Description | abP7Name | 0x0348 | char[8] | 4 | RW | Name of Parameter 7 (filled with spaces) | 840
| abP7Unit | 0x034C | char[8] | 4 | RW | Unit of Parameter 7 (filled with spaces) | 844
| xP7UpperLimit | 0x0350 | float32 | 2 | RW | Upper measuring range of Parameter 7 | 848
| xP7LowerLimit | 0x0352 | float32 | 2 | RW | Lower measuring range of Parameter 7 | 850
| ??? | 0x0354 | ? | 2 |  |  | 852 | 0x9800 0x0000
| ??? | 0x0356 | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 854
| ??? | 0x035E | float32 | 4 |  |  | 862 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x0362 | int16 | 1 |  |  | 866 | 0x0003
| ??? | 0x0363 | int16 | 1 |  |  | 867 | 0x0002
| ??? | 0x0364 | float32 | 8 |  |  | 868 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x036C |  | 84 |  |  | 876 | All zeros
Parameter 8 Description | abP8Name | 0x03C0 | char[8] | 4 | RW | Name of Parameter 8 (filled with spaces) | 960
| abP8Unit | 0x03C4 | char[8] | 4 | RW | Unit of Parameter 8 (filled with spaces) | 964
| xP8UpperLimit | 0x03C8 | float32 | 2 | RW | Upper measuring range of Parameter 8 | 968
| xP8LowerLimit | 0x03CA | float32 | 2 | RW | Lower measuring range of Parameter 8 | 970
| ??? | 0x03CC | ? | 2 |  |  | 972 | 0x9C00 0x0000
| ??? | 0x03CE | float32 | 8 | RW | Calibration coefficients 1-4 Offset (x0), slope (x1), x2, and x3 of local calibration. If global calibration used, will be 0,1,0,0 | 974
| ??? | 0x03D6 | float32 | 4 |  |  | 982 | All not-a-number (0xFFFF or 0x7FC0)
|  | 0x03DA | int16 | 1 |  |  | 986 | 0x0003
| ??? | 0x03DB | int16 | 1 |  |  | 987 | 0x0002
| ??? | 0x03DC | float32 | 8 |  |  | 988 | All not-a-number (0xFFFF or 0x7FC0)
| ??? | 0x03E4 |  | 84 |  |  | 996 | All zeros
Ana::gate repeats the patern of 120 registers/parameter for a total of 32 parameters - ie, until register 3959
Registers 3960 - 3999 are illegal in ana::gate
From register 4000, ana::gate begins to mirror the input registers - continues until register 5001 (equivaled to 1001, the last legal input register in ana::gate)
Global Calibrations | Unknown | 0x0438 | ? | 1 |  |  | 1080 | Actually begins at place pointed to by "pDeviceConfigPrivate"
| cGlobalCal0 | 0x0439 | char[12] | 6 |  | Name of in-use global calibration | 1081
| Unknown | 0x043F | int16 | 1 |  | Index of in use global calibration?? | 1087
| Unknown | 0x0440 | int16 | 1 |  | Number of global cals in the list?? | 1088
| cGlobalCalList1 | 0x0441 | char[12] | 6 |  | Name of next global calibration in storage | 1089 | Name of "GCHECK" ???
| Unknown | 0x0447 | 3 | 1 |  |  | 1095 | Has values 0x0000, 0x00F7 (247), 0x0000
| cGlobalCalList | 0x044A | char | 72 |  | Text list of remaining global calibrations in storage | 1098 | Has just enough space for all of them (spec has room for 12), but doesn't actually list more than 5 for me
Empty and Illegal | Unknown | 0x0492 |  | 2 |  |  | 1170 | All zeros
| Unknown | 0x0494 | char | 17 |  |  | 1172
| Illegal Registers | 0x04A5 | - | 3 |  |  | 1189 | Cannot be read
| Unknown | 0x04A8 | char | 16 |  |  | 1192
| Unknown | 0x04B8 | ? | 4 |  |  | 1208 | Has numbers, but they mean nothing to me
| Unknown | 0x04BC | - | 80 |  |  | 1212 | All zeros
| Illegal Registers | 0x050C | - | 196 |  |  | 1292 | Cannot be read
in-use reference | uiValuesinRef | 0x05D0 | int16 | 1 |  | Number of values recorded for each reference?? | 1488 | ?? Has value 0x0100 = 256
| Unknown | 0x05D1 |  | 18 |  |  | 1489
| uiRefNuminUse | 0x05E3 | int16 | 1 |  | Number of reference in use (0-6) | 1507
| uiRefNameinUse | 0x05E4 | char[8] | 4 |  | Name of reference in use | 1508
| tRefinUse | 0x05E8 | timestamp (TAI64N) | 6 |  | Time reference in use recorded | 1512
Reference 0 | Unknown | 0x05EE |  | 1 |  |  | 1518
| cRefName0 | 0x05EF | char[8] | 4 |  | Name of reference 0 | 1519
| fDarkNoise0 | 0x05F3 | float32 | 2 |  | Dark noise at reference 0 measurement | 1523
| uiAvgK0 | 0x05F5 | int16 | 1 |  | Average K at reference 0 measurement | 1525
| uiAvgM0 | 0x05F6 | int16 | 1 |  | Average M at reference 0 measurement | 1526
| uiFlash0 | 0x05F7 | int16 | 1 |  | Flash Rate [Hz] during reference 0 measurement | 1527
| uiLampV0 | 0x05F8 | int16 | 1 |  | Lamp voltage during reference 0 measurement | 1528
| eDetectorType0 | 0x05F9 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 1529
|  | 0x05FA | int16 | 1 |  | Number of max repetitions | 1530 | What is this?
| eLpfilter0 | 0x05FB | enum | 1 |  | Lp Filter 0 … off 1 … on | 1531
| uiFUG0 | 0x05FC | int16 | 1 |  | Frequency lower limit [Hz] | 1532
| eRefType0 | 0x05FD | enum | 1 |  | Reference type (enum values unknown) | 1533
| eRefOffset0 | 0x05FE | int16 | 1 |  | Reference 0 Offset [abs/m] | 1534
| Unknown | 0x05FF |  | 1 |  |  | 1535
| tRefTime0 | 0x0600 | timestamp (TAI64N) | 6 |  | Time of reference 0 | 1536
| fRefValues0 | 0x0606 | float32 | 512 |  | Raw Absorbances of Reference 0 (186.3-732.7) | 1542 |  | 536
Reference 1 | Unknown | 0x0806 |  | 1 |  |  | 2054
| cRefName1 | 0x0807 | char[8] | 4 |  | Name of Reference 1 | 2055
| fDarkNoise1 | 0x080B | float32 | 2 |  | Dark noise at Reference 1 measurement | 2059
| uiAvgK1 | 0x080D | int16 | 1 |  | Average K at Reference 1 measurement | 2061
| uiAvgM1 | 0x080E | int16 | 1 |  | Average M at Reference 1 measurement | 2062
| uiFlash1 | 0x080F | int16 | 1 |  | Flash Rate [Hz] during Reference 1 measurement | 2063
| uiLampV1 | 0x0810 | int16 | 1 |  | Lamp voltage during Reference 1 measurement | 2064
| eDetectorType1 | 0x0811 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 2065
|  | 0x0812 | int16 | 1 |  | Number of max repetitions | 2066 | What is this?
| eLpfilter1 | 0x0813 | enum | 1 |  | Lp Filter 0 … off 1 … on | 2067
| uiFUG1 | 0x0814 | int16 | 1 |  | Frequency lower limit [Hz] | 2068
| eRefType1 | 0x0815 | enum | 1 |  | Reference type (enum values unknown) | 2069
| eRefOffset1 | 0x0816 | int16 | 1 |  | Reference 1 Offset [abs/m] | 2070
| Unknown | 0x0817 |  | 1 |  |  | 2071
| tRefTime1 | 0x0818 | timestamp (TAI64N) | 6 |  | Time of Reference 1 | 2072
| fRefValues1 | 0x081E | float32 | 512 |  | Raw Absorbances of Reference 1 (186.3-732.7) | 2078
Reference 2 | Unknown | 0x0A1E |  | 1 |  |  | 2590
| cRefName2 | 0x0A1F | char[8] | 4 |  | Name of Reference 2 | 2591
| fDarkNoise2 | 0x0A23 | float32 | 2 |  | Dark noise at Reference 2 measurement | 2595
| uiAvgK2 | 0x0A25 | int16 | 1 |  | Average K at Reference 2 measurement | 2597
| uiAvgM2 | 0x0A26 | int16 | 1 |  | Average M at Reference 2 measurement | 2598
| uiFlash2 | 0x0A27 | int16 | 1 |  | Flash Rate [Hz] during Reference 2 measurement | 2599
| uiLampV2 | 0x0A28 | int16 | 1 |  | Lamp voltage during Reference 2 measurement | 2600
| eDetectorType2 | 0x0A29 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 2601
|  | 0x0A2A | int16 | 1 |  | Number of max repetitions | 2602 | What is this?
| eLpfilter2 | 0x0A2B | enum | 1 |  | Lp Filter 0 … off 1 … on | 2603
| uiFUG2 | 0x0A2C | int16 | 1 |  | Frequency lower limit [Hz] | 2604
| eRefType2 | 0x0A2D | enum | 1 |  | Reference type (enum values unknown) | 2605
| eRefOffset2 | 0x0A2E | int16 | 1 |  | Reference 2 Offset [abs/m] | 2606
| Unknown | 0x0A2F |  | 1 |  |  | 2607
| tRefTime2 | 0x0A30 | timestamp (TAI64N) | 6 |  | Time of Reference 2 | 2608
| fRefValues2 | 0x0A36 | float32 | 512 |  | Raw Absorbances of Reference 2 (186.3-732.7) | 2614
Reference 3 | Unknown | 0x0C36 |  | 1 |  |  | 3126
| cRefName3 | 0x0C37 | char[8] | 4 |  | Name of Reference 3 | 3127
| fDarkNoise3 | 0x0C3B | float32 | 2 |  | Dark noise at Reference 3 measurement | 3131
| uiAvgK3 | 0x0C3D | int16 | 1 |  | Average K at Reference 3 measurement | 3133
| uiAvgM3 | 0x0C3E | int16 | 1 |  | Average M at Reference 3 measurement | 3134
| uiFlash3 | 0x0C3F | int16 | 1 |  | Flash Rate [Hz] during Reference 3 measurement | 3135
| uiLampV3 | 0x0C40 | int16 | 1 |  | Lamp voltage during Reference 3 measurement | 3136
| eDetectorType3 | 0x0C41 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 3137
|  | 0x0C42 | int16 | 1 |  | Number of max repetitions | 3138 | What is this?
| eLpfilter3 | 0x0C43 | enum | 1 |  | Lp Filter 0 … off 1 … on | 3139
| uiFUG3 | 0x0C44 | int16 | 1 |  | Frequency lower limit [Hz] | 3140
| eRefType3 | 0x0C45 | enum | 1 |  | Reference type (enum values unknown) | 3141
| eRefOffset3 | 0x0C46 | int16 | 1 |  | Reference 3 Offset [abs/m] | 3142
| Unknown | 0x0C47 |  | 1 |  |  | 3143
| tRefTime3 | 0x0C48 | timestamp (TAI64N) | 6 |  | Time of Reference 3 | 3144
| fRefValues3 | 0x0C4E | float32 | 512 |  | Raw Absorbances of Reference 3 (186.3-732.7) | 3150
Reference 4 | Unknown | 0x0E4E |  | 1 |  |  | 3662
| cRefName4 | 0x0E4F | char[8] | 4 |  | Name of Reference 4 | 3663
| fDarkNoise4 | 0x0E53 | float32 | 2 |  | Dark noise at Reference 4 measurement | 3667
| uiAvgK4 | 0x0E55 | int16 | 1 |  | Average K at Reference 4 measurement | 3669
| uiAvgM4 | 0x0E56 | int16 | 1 |  | Average M at Reference 4 measurement | 3670
| uiFlash4 | 0x0E57 | int16 | 1 |  | Flash Rate [Hz] during Reference 4 measurement | 3671
| uiLampV4 | 0x0E58 | int16 | 1 |  | Lamp voltage during Reference 4 measurement | 3672
| eDetectorType4 | 0x0E59 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 3673
|  | 0x0E5A | int16 | 1 |  | Number of max repetitions | 3674 | What is this?
| eLpfilter4 | 0x0E5B | enum | 1 |  | Lp Filter 0 … off 1 … on | 3675
| uiFUG4 | 0x0E5C | int16 | 1 |  | Frequency lower limit [Hz] | 3676
| eRefType4 | 0x0E5D | enum | 1 |  | Reference type (enum values unknown) | 3677
| eRefOffset4 | 0x0E5E | int16 | 1 |  | Reference 4 Offset [abs/m] | 3678
| Unknown | 0x0E5F |  | 1 |  |  | 3679
| tRefTime4 | 0x0E60 | timestamp (TAI64N) | 6 |  | Time of Reference 4 | 3680
| fRefValues4 | 0x0E66 | float32 | 512 |  | Raw Absorbances of Reference 4 (186.3-732.7) | 3686
Reference 5 | Unknown | 0x1066 |  | 1 |  |  | 4198
| cRefName5 | 0x1067 | char[8] | 4 |  | Name of Reference 5 | 4199
| fDarkNoise5 | 0x106B | float32 | 2 |  | Dark noise at Reference 5 measurement | 4203
| uiAvgK5 | 0x106D | int16 | 1 |  | Average K at Reference 5 measurement | 4205
| uiAvgM5 | 0x106E | int16 | 1 |  | Average M at Reference 5 measurement | 4206
| uiFlash5 | 0x106F | int16 | 1 |  | Flash Rate [Hz] during Reference 5 measurement | 4207
| uiLampV5 | 0x1070 | int16 | 1 |  | Lamp voltage during Reference 5 measurement | 4208
| eDetectorType5 | 0x1071 | enum | 1 |  | Detector Type 0 … Vis 1 … UV-Vis | 4209
|  | 0x1072 | int16 | 1 |  | Number of max repetitions (??) | 4210
| eLpfilter5 | 0x1073 | enum | 1 |  | Lp Filter 0 … off 1 … on | 4211
| uiFUG5 | 0x1074 | int16 | 1 |  | Frequency lower limit [Hz] | 4212
| eRefType5 | 0x1075 | enum | 1 |  | Reference type (enum values unknown) | 4213
| eRefOffset5 | 0x1076 | int16 | 1 |  | Reference 5 Offset [abs/m] | 4214
| Unknown | 0x1077 |  | 1 |  |  | 4215
| tRefTime5 | 0x1078 | timestamp (TAI64N) | 6 |  | Time of Reference 5 | 4216
| fRefValues5 | 0x107E | float32 | 512 |  | Raw Absorbances of Reference 5 (186.3-732.7) | 4222
Unknown and Illegal | Unknown | 0x127E |  | 2 |  |  | 4734
|  | 0x1280 | int16 | 1 |  |  | 4736
|  | 0x1281 | int16 | 1 |  |  | 4737
|  | 0x1282 | int16 | 1 |  |  | 4738
|  | 0x1283 | int16 | 1 |  |  | 4739
|  | 0x1284 | int16 | 1 |  |  | 4740
|  | 0x1285 | int16 | 1 |  |  | 4741
|  | 0x1286 | int16 | 1 |  |  | 4742
|  | 0x1287 | int16 | 1 |  |  | 4743
|  | 0x1288 | float32 | 2 |  |  | 4744
|  | 0x128A |  |  |  |  | 4746
|  | 0x128C |  |  |  |  | 4748
|  | 0x128E | float32 | 2 |  |  | 4750
|  | 0x1290 |  | 1 |  |  | 4752
|  | 0x1291 |  | 1 |  |  | 4753
|  | 0x1292 |  | 1 |  |  | 4754
|  | 0x1293 |  | 1 |  |  | 4755
|  | 0x1294 |  |  |  |  | 4756
|  | 0x1296 |  |  |  |  | 4758
|  | 0x1298 |  |  |  |  | 4760
|  | 0x129A |  | 1 |  |  | 4762
| Unknown | 0x129B |  | 1 |  |  | 4763
| Illegal Registers | 0x129C | - | 2 |  |  | 4764
| Unknown | 0x129E |  | 70 |  |  | 4766 | All zeros
| Illegal Registers | 0x12E4 |  | - |  |  | 4836 | All remaining registers are illegal

## Input Registers
Spectro::lyzer Input Registers
Read Command = 0x04
All Registers are Read-Only
Group | Tag Name | Register | Type | Length | R/W | Description | Address | Notes
device id. | uiVersion | 0x0000 | unit16 | 1 | R | Version of Modbus mapping protocol. For all changes in public registers: 0xAABB AA ... Major version BB ... Minor version -compatible | 0
| eVendor | 0x0001 | enum | 1 | R | Vendor code 0x0000 … unknown 0x96C3 … s::can | 1
| eModel | 0x0002 | enum | 1 | R | Device model 0x0000 - unknown 0x0101 - spectro::lyzer 0x0603 - con::stat/con::cube/ana::gate | 2
device description | abModel | 0x0003 | char[20] | 10 | R | Description of device model, filled with spaces (or 0's) | 3
| abSerialNumber | 0x000D | char[8] | 4 | R | Serial number, filled with spaces (or 0's) | 13
| abHWRelease | 0x0011 | char[4] | 2 | R | Hardware release: 0xAABB AA ... Major version BB ... Minor version | 17
| abSWRelease | 0x0013 | char[4] | 2 | R | Software release: 0xAABB AA ... Major version BB ... Minor version | 19
| uiHWStarts | 0x0015 | unit16 | 1 | R | Device rebooter counter | 21
| uiParameterCount | 0x0016 | unit16 | 1 | R | Number of Parameters | 22
| eParameterType | 0x0017 | enum | 1 | R | Data type of Parameter and Parameter limits (check for compatibility) | 23
| uiParameterScale | 0x0018 | unit16 | 1 | R | Parameter scale factor. Used for all Parameter values which depend on eParameterType. | 24
Unknows | Unknown | 0x0019 |  | 6 | R |  | 25 | All zeros
| Unknown | 0x001F |  | 1 | R |  | 31 | Has a value…
| Unknown | 0x0020 | char[20] | 10 | R |  | 32 | A ascii character string with a date.. Unknown what the date is for (warranty expiration date? firmware update date? gate last factory serviced?). Spectro::lyzer only (not ana::gate)
| Unknown | 0x002A |  | 62 |  |  | 42 | All zeros
device status public | tSampleTime | 0x0068 | timestamp (TAI64N) | 6 | R | Time when the Parameter results have been updated. Timestamp of logged status and results | 104
| Unknown | 0x006E |  | 10 |  |  | 110 | All zeros
| bmDeviceStatus | 0x0078 | bitmask | 1 |  | Device Status ... See bitmask map | 120
| Unknown | 0x0079 |  | 1 |  |  | 121
| Unknown | 0x007A |  | 1 |  | WAG = May be a pointer to the first register with fingerprint data | 122 | Has value 0x01FD=509 from both spec and ana::gate
| Unknown | 0x007B |  | 5 |  |  | 123 | All zeros
Parameter 1 Result | bmP1Status | 0x0080 | bitmask | 1 | R | Parameter 1 status ... See bitmask map | 128
| bmP1PrivStatus | 0x0081 | bitmask | 1 | R | Parameter 1 private status | 129
| xP1Value | 0x0082 | float32 | 2 | R | Parameter 1 result | 130
| xP1PrivValue | 0x0084 | float32 | 2 | R | Parameter 1 result private | 132
| Reserved | 0x0086 | - | 2 | R |  | 134
Parameter 2 Result | bmP2Status | 0x0088 | bitmask | 1 | R | Parameter 2 status ... See bitmask map | 136
| bmP2PrivStatus | 0x0089 | bitmask | 1 | R | Parameter 2 private status | 137
| xP2Value | 0x008A | float32 | 2 | R | Parameter 2 result | 138
| xP2PrivValue | 0x008C | float32 | 2 | R | Parameter 2 result private | 140
| Reserved | 0x008E | - | 2 | R |  | 142
Parameter 3 Result | bmP3Status | 0x0090 | bitmask | 1 | R | Parameter 3 status ... See bitmask map | 144
| bmP3PrivStatus | 0x0091 | bitmask | 1 | R | Parameter 3 private status | 145
| xP3Value | 0x0092 | float32 | 2 | R | Parameter 3 result | 146
| xP3PrivValue | 0x0094 | float32 | 2 | R | Parameter 3 result private | 148
| Reserved | 0x0096 | - | 2 | R |  | 150
Parameter 4 Result | bmP4Status | 0x0098 | bitmask | 1 | R | Parameter 4 status ... See bitmask map | 152
| bmP4PrivStatus | 0x0099 | bitmask | 1 | R | Parameter 4 private status | 153
| xP4Value | 0x009A | float32 | 2 | R | Parameter 4 result | 154
| xP4PrivValue | 0x009C | float32 | 2 | R | Parameter 4 result private | 156
| Reserved | 0x009E | - | 2 | R |  | 158
Parameter 5 Result | bmP5Status | 0x00A0 | bitmask | 1 | R | Parameter 5 status ... See bitmask map | 160
| bmP5PrivStatus | 0x00A1 | bitmask | 1 | R | Parameter 5 private status | 161
| xP5Value | 0x00A2 | float32 | 2 | R | Parameter 5 result | 162
| xP5PrivValue | 0x00A4 | float32 | 2 | R | Parameter 5 result private | 164
| Reserved | 0x00A6 | - | 2 | R |  | 166
Parameter 6 Result | bmP6Status | 0x00A8 | bitmask | 1 | R | Parameter 6 status ... See bitmask map | 168
| bmP6PrivStatus | 0x00A9 | bitmask | 1 | R | Parameter 6 private status | 169
| xP6Value | 0x00AA | float32 | 2 | R | Parameter 6 result | 170
| xP6PrivValue | 0x00AC | float32 | 2 | R | Parameter 6 result private | 172
| Reserved | 0x00AE | - | 2 | R |  | 174
Parameter 7 Result | bmP7Status | 0x00B0 | bitmask | 1 | R | Parameter 7 status ... See bitmask map | 176
| bmP7PrivStatus | 0x00B1 | bitmask | 1 | R | Parameter 7 private status | 177
| xP7Value | 0x00B2 | float32 | 2 | R | Parameter 7 result | 178
| xP7PrivValue | 0x00B4 | float32 | 2 | R | Parameter 7 result private | 180
| Reserved | 0x00B6 | - | 2 | R |  | 182
Parameter 8 Result | bmP8Status | 0x00B8 | bitmask | 1 | R | Parameter 8 status ... See bitmask map | 184
| bmP8PrivStatus | 0x00B9 | bitmask | 1 | R | Parameter 8 private status | 185
| xP8Value | 0x00BA | float32 | 2 | R | Parameter 8 result | 186
| xP8PrivValue | 0x00BC | float32 | 2 | R | Parameter 8 result private | 188
| Reserved | 0x00BE | - | 2 | R |  | 190
Parameter 9 Result | bmP9Status | 0x00C0 | bitmask | 1 | R | Parameter 9 status ... See bitmask map | 192 | ana:: gate only
| bmP9PrivStatus | 0x00C1 | bitmask | 1 | R | Parameter 9 private status | 193
| xP9Value | 0x00C2 | float32 | 2 | R | Parameter 9 result | 194
| xP9PrivValue | 0x00C4 | float32 | 2 | R | Parameter 9 result private | 196
| Reserved | 0x00C6 | - | 2 | R |  | 198
Parameter 10 Result | bmP10Status | 0x00C8 | bitmask | 1 | R | Parameter 10 status ... See bitmask map | 200
| bmP10PrivStatus | 0x00C9 | bitmask | 1 | R | Parameter 10 private status | 201
| xP10Value | 0x00CA | float32 | 2 | R | Parameter 10 result | 202
| xP10PrivValue | 0x00CC | float32 | 2 | R | Parameter 10 result private | 204
| Reserved | 0x00CE | - | 2 | R |  | 206
Parameter 11 Result | bmP11Status | 0x00D0 | bitmask | 1 | R | Parameter 11 status ... See bitmask map | 208
| bmP11PrivStatus | 0x00D1 | bitmask | 1 | R | Parameter 11 private status | 209
| xP11Value | 0x00D2 | float32 | 2 | R | Parameter 11 result | 210
| xP11PrivValue | 0x00D4 | float32 | 2 | R | Parameter 11 result private | 212
| Reserved | 0x00D6 | - | 2 | R |  | 214
Parameter 12 Result | bmP12Status | 0x00D8 | bitmask | 1 | R | Parameter 12 status ... See bitmask map | 216
| bmP12PrivStatus | 0x00D9 | bitmask | 1 | R | Parameter 12 private status | 217
| xP12Value | 0x00DA | float32 | 2 | R | Parameter 12 result | 218
| xP12PrivValue | 0x00DC | float32 | 2 | R | Parameter 12 result private | 220
| Reserved | 0x00DE | - | 2 | R |  | 222
Parameter 13 Result | bmP13Status | 0x00E0 | bitmask | 1 | R | Parameter 13 status ... See bitmask map | 224
| bmP13PrivStatus | 0x00E1 | bitmask | 1 | R | Parameter 13 private status | 225
| xP13Value | 0x00E2 | float32 | 2 | R | Parameter 13 result | 226
| xP13PrivValue | 0x00E4 | float32 | 2 | R | Parameter 13 result private | 228
| Reserved | 0x00E6 | - | 2 | R |  | 230
Parameter 14 Result | bmP14Status | 0x00E8 | bitmask | 1 | R | Parameter 14 status ... See bitmask map | 232
| bmP14PrivStatus | 0x00E9 | bitmask | 1 | R | Parameter 14 private status | 233
| xP14Value | 0x00EA | float32 | 2 | R | Parameter 14 result | 234
| xP14PrivValue | 0x00EC | float32 | 2 | R | Parameter 14 result private | 236
| Reserved | 0x00EE | - | 2 | R |  | 238
Parameter 15 Result | bmP15Status | 0x00F0 | bitmask | 1 | R | Parameter 15 status ... See bitmask map | 240
| bmP15PrivStatus | 0x00F1 | bitmask | 1 | R | Parameter 15 private status | 241
| xP15Value | 0x00F2 | float32 | 2 | R | Parameter 15 result | 242
| xP15PrivValue | 0x00F4 | float32 | 2 | R | Parameter 15 result private | 244
| Reserved | 0x00F6 | - | 2 | R |  | 246
Parameter 16 Result | bmP16Status | 0x00F8 | bitmask | 1 | R | Parameter 16 status ... See bitmask map | 248
| bmP16PrivStatus | 0x00F9 | bitmask | 1 | R | Parameter 16 private status | 249
| xP16Value | 0x00FA | float32 | 2 | R | Parameter 16 result | 250
| xP16PrivValue | 0x00FC | float32 | 2 | R | Parameter 16 result private | 252
| Reserved | 0x00FE | - | 2 | R |  | 254
Parameter 17 Result | bmP17Status | 0x0100 | bitmask | 1 | R | Parameter 17 status ... See bitmask map | 256
| bmP17PrivStatus | 0x0101 | bitmask | 1 | R | Parameter 17 private status | 257
| xP17Value | 0x0102 | float32 | 2 | R | Parameter 17 result | 258
| xP17PrivValue | 0x0104 | float32 | 2 | R | Parameter 17 result private | 260
| Reserved | 0x0106 | - | 2 | R |  | 262
Parameter 18 Result | bmP18Status | 0x0108 | bitmask | 1 | R | Parameter 18 status ... See bitmask map | 264
| bmP18PrivStatus | 0x0109 | bitmask | 1 | R | Parameter 18 private status | 265
| xP18Value | 0x010A | float32 | 2 | R | Parameter 18 result | 266
| xP18PrivValue | 0x010C | float32 | 2 | R | Parameter 18 result private | 268
| Reserved | 0x010E | - | 2 | R |  | 270
Parameter 19 Result | bmP19Status | 0x0110 | bitmask | 1 | R | Parameter 19 status ... See bitmask map | 272
| bmP19PrivStatus | 0x0111 | bitmask | 1 | R | Parameter 19 private status | 273
| xP19Value | 0x0112 | float32 | 2 | R | Parameter 19 result | 274
| xP19PrivValue | 0x0114 | float32 | 2 | R | Parameter 19 result private | 276
| Reserved | 0x0116 | - | 2 | R |  | 278
Parameter 20 Result | bmP20Status | 0x0118 | bitmask | 1 | R | Parameter 20 status ... See bitmask map | 280
| bmP20PrivStatus | 0x0119 | bitmask | 1 | R | Parameter 20 private status | 281
| xP20Value | 0x011A | float32 | 2 | R | Parameter 20 result | 282
| xP20PrivValue | 0x011C | float32 | 2 | R | Parameter 20 result private | 284
| Reserved | 0x011E | - | 2 | R |  | 286
Parameter 21 Result | bmP21Status | 0x0120 | bitmask | 1 | R | Parameter 21 status ... See bitmask map | 288
| bmP21PrivStatus | 0x0121 | bitmask | 1 | R | Parameter 21 private status | 289
| xP21Value | 0x0122 | float32 | 2 | R | Parameter 21 result | 290
| xP21PrivValue | 0x0124 | float32 | 2 | R | Parameter 21 result private | 292
| Reserved | 0x0126 | - | 2 | R |  | 294
Parameter 22 Result | bmP22Status | 0x0128 | bitmask | 1 | R | Parameter 22 status ... See bitmask map | 296
| bmP22PrivStatus | 0x0129 | bitmask | 1 | R | Parameter 22 private status | 297
| xP22Value | 0x012A | float32 | 2 | R | Parameter 22 result | 298
| xP22PrivValue | 0x012C | float32 | 2 | R | Parameter 22 result private | 300
| Reserved | 0x012E | - | 2 | R |  | 302
Parameter 23 Result | bmP23Status | 0x0130 | bitmask | 1 | R | Parameter 23 status ... See bitmask map | 304
| bmP23PrivStatus | 0x0131 | bitmask | 1 | R | Parameter 23 private status | 305
| xP23Value | 0x0132 | float32 | 2 | R | Parameter 23 result | 306
| xP23PrivValue | 0x0134 | float32 | 2 | R | Parameter 23 result private | 308
| Reserved | 0x0136 | - | 2 | R |  | 310
Parameter 24 Result | bmP24Status | 0x0138 | bitmask | 1 | R | Parameter 24 status ... See bitmask map | 312
| bmP24PrivStatus | 0x0139 | bitmask | 1 | R | Parameter 24 private status | 313
| xP24Value | 0x013A | float32 | 2 | R | Parameter 24 result | 314
| xP24PrivValue | 0x013C | float32 | 2 | R | Parameter 24 result private | 316
| Reserved | 0x013E | - | 2 | R |  | 318
Parameter 25 Result | bmP25Status | 0x0140 | bitmask | 1 | R | Parameter 25 status ... See bitmask map | 320
| bmP25PrivStatus | 0x0141 | bitmask | 1 | R | Parameter 25 private status | 321
| xP25Value | 0x0142 | float32 | 2 | R | Parameter 25 result | 322
| xP25PrivValue | 0x0144 | float32 | 2 | R | Parameter 25 result private | 324
| Reserved | 0x0146 | - | 2 | R |  | 326
Parameter 26 Result | bmP26Status | 0x0148 | bitmask | 1 | R | Parameter 26 status ... See bitmask map | 328
| bmP26PrivStatus | 0x0149 | bitmask | 1 | R | Parameter 26 private status | 329
| xP26Value | 0x014A | float32 | 2 | R | Parameter 26 result | 330
| xP26PrivValue | 0x014C | float32 | 2 | R | Parameter 26 result private | 332
| Reserved | 0x014E | - | 2 | R |  | 334
Parameter 27 Result | bmP27Status | 0x0150 | bitmask | 1 | R | Parameter 27 status ... See bitmask map | 336
| bmP27PrivStatus | 0x0151 | bitmask | 1 | R | Parameter 27 private status | 337
| xP27Value | 0x0152 | float32 | 2 | R | Parameter 27 result | 338
| xP27PrivValue | 0x0154 | float32 | 2 | R | Parameter 27 result private | 340
| Reserved | 0x0156 | - | 2 | R |  | 342
Parameter 28 Result | bmP28Status | 0x0158 | bitmask | 1 | R | Parameter 28 status ... See bitmask map | 344
| bmP28PrivStatus | 0x0159 | bitmask | 1 | R | Parameter 28 private status | 345
| xP28Value | 0x015A | float32 | 2 | R | Parameter 28 result | 346
| xP28PrivValue | 0x015C | float32 | 2 | R | Parameter 28 result private | 348
| Reserved | 0x015E | - | 2 | R |  | 350
Parameter 29 Result | bmP29Status | 0x0160 | bitmask | 1 | R | Parameter 29 status ... See bitmask map | 352
| bmP29PrivStatus | 0x0161 | bitmask | 1 | R | Parameter 29 private status | 353
| xP29Value | 0x0162 | float32 | 2 | R | Parameter 29 result | 354
| xP29PrivValue | 0x0164 | float32 | 2 | R | Parameter 29 result private | 356
| Reserved | 0x0166 | - | 2 | R |  | 358
Parameter 30 Result | bmP30Status | 0x0168 | bitmask | 1 | R | Parameter 30 status ... See bitmask map | 360
| bmP30PrivStatus | 0x0169 | bitmask | 1 | R | Parameter 30 private status | 361
| xP30Value | 0x016A | float32 | 2 | R | Parameter 30 result | 362
| xP30PrivValue | 0x016C | float32 | 2 | R | Parameter 30 result private | 364
| Reserved | 0x016E | - | 2 | R |  | 366
Parameter 31 Result | bmP31Status | 0x0170 | bitmask | 1 | R | Parameter 31 status ... See bitmask map | 368
| bmP31PrivStatus | 0x0171 | bitmask | 1 | R | Parameter 31 private status | 369
| xP31Value | 0x0172 | float32 | 2 | R | Parameter 31 result | 370
| xP31PrivValue | 0x0174 | float32 | 2 | R | Parameter 31 result private | 372
| Reserved | 0x0176 | - | 2 | R |  | 374
Parameter 32 Result | bmP32Status | 0x0178 | bitmask | 1 | R | Parameter 32 status ... See bitmask map | 376
| bmP32PrivStatus | 0x0179 | bitmask | 1 | R | Parameter 32 private status | 377
| xP32Value | 0x017A | float32 | 2 | R | Parameter 32 result | 378
| xP32PrivValue | 0x017C | float32 | 2 | R | Parameter 32 result private | 380
| Reserved | 0x017E | - | 2 | R |  | 382
| Illegal Registers | 0x00C0 / 0x0180 | - | 316 / 124 | R |  | 192 / 384 | Cannot be read. Illegal registers begin at register 192 for spectro::lyzer
Fingerprint (or Selected Fingerprint type) | Unknown | 0x01FC |  | 1 | R |  | 508 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x01FD |  | 3 | R |  | 509 | All zeros
| tFingerprintTime | 0x0200 | timestamp (TAI64N) | 6 | R |  | 512
| eDetectorType | 0x0206 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 518 | Guessing
| eSpectralSource | 0x0207 | enum | 1 |  | Spectral Source 0 … Fingerprint [Abs/m] 1 … Turbidity compensated fingerprint [Abs/m] 2 … First derivative of fingerprint [Abs/m/m] 3 … Difference between current print and print in memory [Abs/m] 4 … Percent Transmission [%/cm2] 5 … First derivative of turbidity compensated fingerprint [Abs/m/m] 6 ... Percent Transmission per 10cm2 [%/10cm2] 7 ... UNKNOWN | 519 | Guessing
| uiPathLength | 0x0208 | int16 | 1 |  | Spectral Path Length | 520 | Guessing
| unknown | 0x0209 |  | 1 |  | ?? Fp status?? | 521
| fFingerprintData | 0x020A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 522
| unknown | 0x03C4 |  | 56 |  |  | 964 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Ana::gate lists the sensor global calibration as ascii characteters in registers 964 - 969
Ana::gate has 0x0001 in 970 (type of global cal??) then 0x0101 in 971 (??)
Ana::gate lists the sensor name fro 972-986
Ana::gate lists the station name from 987 - 1001
Registers above 1001 are illegal in ana::gate
Turbidity Compensated Fingerprint [abs/m] | Unknown | 0x03FC |  | 1 | R | WAG = Maybe the number of values in the fingerprint | 1020 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x03FD |  | 3 | R |  | 1021 | All zeros
| tFingerprintTime | 0x0400 | timestamp (TAI64N) | 6 | R |  | 1024
| eDetectorType | 0x0406 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 1030 | Guessing
| eSpectralSource | 0x0407 | enum | 1 |  | Spectral Source 1 … Turbidity compensated fingerprint [Abs/m] | 1031 | Guessing
| uiPathLength | 0x0408 | int16 | 1 |  | Spectral Path Length | 1032 | Guessing
| unknown | 0x0409 |  | 1 |  | ?? Fp status?? | 1033
| fFingerprintData | 0x040A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 1034
| unknown | 0x05C4 |  | 56 |  |  | 1476 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
First derivative of the measured fingerprint (i.e. gradient)[Abs/m] | Unknown | 0x05FC |  | 1 | R |  | 1532 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x05FD |  | 3 | R |  | 1533 | All zeros
| tFingerprintTime | 0x0600 | timestamp (TAI64N) | 6 | R |  | 1536
| eDetectorType | 0x0606 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 1542 | Guessing
| eSpectralSource | 0x0607 | enum | 1 |  | Spectral Source 2 … First derivative of fingerprint [Abs/m/m] | 1543 | Guessing
| uiPathLength | 0x0608 | int16 | 1 |  | Spectral Path Length | 1544 | Guessing
| unknown | 0x0609 |  | 1 |  | ?? Fp status?? | 1545
| fFingerprintData | 0x060A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 1546
| unknown | 0x07C4 |  | 56 |  |  | 1988 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Difference between the current fingerprint and the previous one in memory [Abs/m] | Unknown | 0x07FC |  | 1 | R |  | 2044 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x07FD |  | 3 | R |  | 2045 | All zeros
| tFingerprintTime | 0x0800 | timestamp (TAI64N) | 6 | R |  | 2048
| eDetectorType | 0x0806 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 2054 | Guessing
| eSpectralSource | 0x0807 | enum | 1 |  | Spectral Source 3 … Difference between current print and print in memory [Abs/m] | 2055 | Guessing
| uiPathLength | 0x0808 | int16 | 1 |  | Spectral Path Length | 2056 | Guessing
| unknown | 0x0809 |  | 1 |  | ?? Fp status?? | 2057
| fFingerprintData | 0x080A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 2058
| unknown | 0x09C4 |  | 56 |  |  | 2500 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Percent transmission - NOT linear wrt concentrations [%/cm2] | Unknown | 0x09FC |  | 1 | R |  | 2556 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x09FD |  | 3 | R |  | 2557 | All zeros
| tFingerprintTime | 0x0A00 | timestamp (TAI64N) | 6 | R |  | 2560
| eDetectorType | 0x0A06 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 2566 | Guessing
| eSpectralSource | 0x0A07 | enum | 1 |  | Spectral Source 4 … Percent Transmission [%/cm2] | 2567 | Guessing
| uiPathLength | 0x0A08 | int16 | 1 |  | Spectral Path Length | 2568 | Guessing
| unknown | 0x0A09 |  | 1 |  | ?? Fp status?? | 2569
| fFingerprintData | 0x0A0A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 2570
| unknown | 0x0BC4 |  | 56 |  |  | 3012 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
First derivative of the turbidity-compensated fingerprint [Abs/m] | Unknown | 0x0BFC |  | 1 | R |  | 3068 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x0BFD |  | 3 | R |  | 3069 | All zeros
| tFingerprintTime | 0x0C00 | timestamp (TAI64N) | 6 | R |  | 3072
| eDetectorType | 0x0C06 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 3078 | Guessing
| eSpectralSource | 0x0C07 | enum | 1 |  | Spectral Source 5 … First derivative of turbidity compensated fingerprint [Abs/m/m] | 3079 | Guessing
| uiPathLength | 0x0C08 | int16 | 1 |  | Spectral Path Length | 3080 | Guessing
| unknown | 0x0C09 |  | 1 |  | ?? Fp status?? | 3081
| fFingerprintData | 0x0C0A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 3082
| unknown | 0x0DC4 |  | 56 |  |  | 3524 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Percent transmission per 10 cm2 [%/10cm2] | Unknown | 0x0DFC |  | 1 | R |  | 3580 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x0DFD |  | 3 | R |  | 3581 | All zeros
| tFingerprintTime | 0x0E00 | timestamp (TAI64N) | 6 | R |  | 3584
| eDetectorType | 0x0E06 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 3590 | Guessing
| eSpectralSource | 0x0E07 | enum | 1 |  | Spectral Source 6 ... Percent Transmission per 10cm2 [%/10cm2] | 3591 | Guessing
| uiPathLength | 0x0E08 | int16 | 1 |  | Spectral Path Length | 3592 | Guessing
| unknown | 0x0E09 |  | 1 |  | ?? Fp status?? | 3593
| fFingerprintData | 0x0E0A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 3594
| unknown | 0x0FC4 |  | 56 |  |  | 4036 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Unknown type of fingerprint data | Unknown | 0x0FFC |  | 1 | R |  | 4092 | Has value of 0x0100=256 for the first fingerprint and 0x0000 for all others
| Unknown | 0x0FFD |  | 3 | R |  | 4093 | All zeros
| tFingerprintTime | 0x1000 | timestamp (TAI64N) | 6 | R |  | 4096
| eDetectorType | 0x1006 | enum | 1 | R | Fingerprint Detector Type 0 … Vis 1 … UV-Vis | 4102 | Guessing
| eSpectralSource | 0x1007 | enum | 1 |  | Spectral Source 7 ... UNKNOWN | 4103 | Guessing
| uiPathLength | 0x1008 | int16 | 1 |  | Spectral Path Length | 4104 | Guessing
| unknown | 0x1009 |  | 1 |  | ?? Fp status?? | 4105
| fFingerprintData | 0x100A | float32 | 442 |  | "Raw" fingerprint [Abs/m] | 4106
| unknown | 0x11C4 |  | 56 |  |  | 4548 | All zeros - I'm guessing this is actually padding in case the fingerprint actually had 256 values
Unknown and Illegal | Unknown |  |  | 18 |  |  | 4604 | All zeros
| Unknown |  |  | 2 |  |  | 4622
| Unknown |  |  | 4 |  |  | 4624 | All remaining registers are illegal
| Illegal Registers | [5 empty] | 4628 | All remaining registers are illegal

## Bitmasks
Device Status
Bit Number | Significance
15 | Device maintenance required
14 | Device cleaning required
13 | Device busy
12
11
10
9
8
7
6
5
4
3 | Data logger error, no readings can be stored because datalogger is full
2 | Missing or devective component detected
1 | Probe misuse, operation outside the specified temperature range
0 | s::can device reports error during internal check
No bits present | Device is operating normally
System Status
Bit Number | Significance
15
14
13
12
11
10
9
8
7
6 | mA signal is outside of the allowed input range
5 | Validation results are not available
4
3
2
1 | Invalid prove/sensor; serial number of probe/sensor is different
0 | No communication between probe/sensor and controller
No bits present | System is operating normally
Parameter Status
Bit Number | Significance
15 | Parameter reading out of measuring range
14 | Status of alarm Parameter is 'WARNING'
13 | Status of alarm Parameter is 'ALARM'
12
11 | Maintenance necessary
10
9
8
7
6
5 | Parameter not ready or not available
4 | Incorrect calibration, at least one calibration coefficient invalid
3 | Parameter error, the sensor is outside of the medium or in incorrect medium
2 | Parameter error, calibration error
1 | Parameter error, hardware error
0 | General parameter error, at least one internal parameter check failed
No bits present | Parameter is operating normally